## Slide 1
A quick how to….
Items in bold are actions. For example, press the next button…

## Slide 2
If Powerpoint gave you security pop-up, you must click enable content for external media objects to fully run in this demo

## Slide 3
A quick how to….
Items that highlight are clickable. Move your mouse over the next button, and press the next button

## Slide 4
A quick how to….
You will see helpful arrows to guide you where to click. Press here.

## Slide 5
A quick how to….
Hyperlinks are clickable. Try it.

## Slide 6
A quick how to….
These buttons are clickable. Hover mouse over buttons for description of action. Press here.

## Slide 7
A quick how to….
Later, you may decide to click here , then arrow options, and make your cursor visible the whole time

## Slide 8
A quick how to….
Now press the Virtual Tour menu button to get started

## Slide 9
Buttons and Connections
Getting Started with the Certifier
XpertCare Support
Reporting Software
4500 Virtual Tour Menu, Copper
Recommend starting here

## Slide 10
Getting started
If this is the first time you’ve ever used your WireXpert, use the 5 step menu on the next slide, and click through each step to get familiar with the WireXpert.
This 5 step menu mimics the quick start guide
Your settings and preferences may vary, but this demo will provide an overview

## Slide 11
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest

## Slide 12
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest
Press System Settings
Press

## Slide 13
Getting started
Personalize your system settings
  What it is:
    Sound on/off, English/Metric, auto power off interval
  Why you need it:
    These are your personal preferences for the certifier. For example, you might want your menu clicks to be silent, or make an audible click when a button is pressed. Or you might like the chime when the remote unit is connected to the local unit.
Getting started – System Settings

## Slide 14
Press M for meters
Now we are going to walk through examples of preferences you might set
First, change unit of length from feet to meters
FT is feet
M is meters
Getting started – System Settings
Press

## Slide 15
One test that is run is called Near End Cross Talk (NEXT)
The results of NEXT are displayed on a Cartesian plot (X vs Y) where the X-axis is frequency and the Y-axis is signal strength
The NEXT graph will always show the test limits as a red curve. The test results should not exceed the test limits.
Explanation of Plot Y-axis Direction:
Getting started – System Settings

## Slide 16
The red curve is the test limit line.
Users may prefer that the test limit line be displayed above the data, or users may prefer that the test limit line be displayed below the data.
Here are the two options:
Red test limit line below data
Red test limit line above data
Explanation of Plot Y-axis Direction:
Getting started – System Settings

## Slide 17
Press switch (on certifier) and change Plot Y-axis Direction
Switch position for red test limit line above
Explanation of Plot Y-axis Direction:
Getting started – System Settings
Press

## Slide 18
Switch position for red test limit line below
Explanation of Plot Y-axis Direction:
Getting started – System Settings
Press

## Slide 19
Press Sleep Interval
Adjust the sleep interval. You may want to do this in certain cases when you are trying to extend battery life
Getting started – System Settings
Press

## Slide 20
Press the check button (on the certifier) to keep the defaults to never sleep and never power off
Getting started – System Settings
Press

## Slide 21
Press Audio Settings
Getting started – System Settings
Press

## Slide 22
From this screen
Turn off audible touch clicks
Turn off speaker and tone
Turn off voice prompts
Adjust volume (0% to 100%)
Keep the defaults and press
Getting started – System Settings
Press

## Slide 23
Press back button to return to the main setup screen
Done with Personalization
Getting started – System Settings
Press

## Slide 24
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest
Press Project Settings
Getting Started – Project Settings
Press

## Slide 25
What it is:
  Create site names
  Create user names
  Create cable names or cable labels
Why you need it:
  Your site name helps you organize your certification results. Think of this as your folder for certification results with the site name as the folder name
  Operator Name helps organize certification results for each user performing certifications. You may have more than one site or user on a certifier
  Label Source is where you choose your method of labeling cables. You may have 1 cable to certify or 1000s of cables to certify. We give you 4 powerful methods to label or name your cables to maximize your efficiency
Getting Started – Project Settings

## Slide 26
Site
This creates a folder for Autotest results
When a user presses the OneTouch Autotest button (Autotest covered at the end of this section) and runs a cable test, the results of the cable test will be placed in this user defined folder
Creating a site folder is a great way to organize data for a given job site
Press Site to create a site
Getting Started – Project Settings
Press
A year later, a customer calls up, says, “I need you to install 20 new cable runs where you installed 500 last year, and I need you to certify them.” No problem. You pull up your stored site folder in the WireXpert. You instantly see who performed the work last, type of cable installed, type of connector used, and in the notes section, you even documented the customer’s cable color preference. Now you have all the information you need to order and show up with the right cable, connectors, and even the right cable color.
Application Story:

## Slide 27
Site
You can select a site from the list of sites previously created (4 items are shown in the list right now)
Or, Press the to create a new site
Press the and make a new site
Getting Started – Project Settings
Press

## Slide 28
Site
Type your site name (Press in the space provided to type New Site Name)
Getting Started – Project Settings
Press

## Slide 29
West Towne Mall
Now press to accept your site name
Site
Type your site name (Press in the space provided to type)
Getting Started – Project Settings
Press
Press

## Slide 30
West Towne Mall
7600 Kingston Pike
Site
Now press to accept your site address and notes (maybe you entered a cable color preference)
Getting Started – Project Settings
Press

## Slide 31
West Towne Mall
7600 Kingston Pike
Site
Now you can see that the site name is West Towne Mall
Now press Operator (below Site) to assign or create a new Operator for this job
Getting Started – Project Settings
Press

## Slide 32
Operator
Let’s setup Mark as the operator, and select Mark from the list of operators
West Towne Mall
7600 Kingston Pike
Getting Started – Project Settings
Press

## Slide 33
Label Source
Now you see site name and the Operator, Mark, is set for Project settings
Next, let’s define how we will label each cable to be tested
Select Label Source
Getting Started – Project Settings
Press

## Slide 34
Label Source
There are 4 convenient options because this is an important part of ease of use…
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Idea for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  Extremely powerful cable label tool where a wizard helps you generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing may not happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings

## Slide 35
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Idea for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  Extremely powerful cable label tool where a wizard helps you generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing may not happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings

## Slide 36
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Ideal for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  Extremely powerful cable label tool where a wizard helps you generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing may not happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings

## Slide 37
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Ideal for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
  Free Form, create multiple outlet faceplates
List-based testing
  Extremely powerful cable label tool where a wizard helps you generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing may not happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings

## Slide 38
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Ideal for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  The most powerful cable label tool, and a key reason cable installers love the WireXpert. Generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing doesn’t happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings

## Slide 39
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  Ideal for just a few cables to be tested with random names (only recommended for just a few cables)
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Ideal for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  The most powerful cable label tool, and a key reason cable installers love the WireXpert. Generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing doesn’t happen in order
Select Simple Cable Label because we are testing a patch panel
Getting Started – Project Settings
Press

## Slide 40
Label Source
There are 4 convenient options because this is an important part of ease of use
None
  Just name each cable manually after saving a result
  ideal for just a few cables to be tested with random names
Simple Cable Label
  Name the cable, then a start number. The cable number will increment by one each time you save a new test
  Idea for a simple cable scheme where cables are in sequential order
Template label
  Templates have been included for you based on international standard labeling schemes, for example, TIA-606-A Class 2 standard labeling scheme
  Ideal if you are required to follow and certify cables using a standards based cable labeling scheme
List-based testing
  Extremely powerful cable label tool where a wizard helps you generate 1000’s of cable labels with just a few mouse clicks
  Ideal for large jobs, with lots of cable labels, and when testing may not happen in order
Select Simple Cable Label because we only have a few cables to test
Getting Started – Project Settings
Press

## Slide 41
Label Source
Enter your cable Label Name
  “A” was typed here for you
  This will be a constant
Enter your Start Value (a cable number) for the first cable to be tested
  “1” was typed here for you
  Each time Autotest (covered at the end of this section) is pressed, this value will increment
Press
Getting Started – Project Settings
Press

## Slide 42
AutoSave test results
AutoSave has a switch with setting of 0 and 1
  0 is AutoSave off
  1 is AutoSave on
  Press 1 to enable AutoSave so that when a test passes, the test result will automatically be saved to your project folder
Press 1 now
Getting Started – Project Settings
Press

## Slide 43
Project settings are now complete
You can see that the first cable to be tested is labeled A-1
Press the back button to return to the main setup menu
Getting Started – Project Settings
Press

## Slide 44
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest
Press Test Settings
Getting Started – Test Settings
Press

## Slide 45
Test Settings
What it is:
  For a given project, you will be required to certify the cables to a specific standard
  Example certification standards are TIA CAT 6 or TIA CAT 6A
Why you need it:
  Test Limit is where you select the required certification standard for the job
  The WireXpert will adapt the test for you depending on the cable you enter
  You don’t need to be a certification expert, you only need to know what kind of cable and connector you have. The WireXpert will figure out the rest.
  In fact, we have hundreds of cable manufacturers and their NVP (Nominal Velocity of Propagation) already entered for you. You only have to select from a list.
Getting Started – Test Settings
Press Test Limit
Press

## Slide 46
Press TIA
to select a list of TIA standards
Getting Started – Test Settings
Press

## Slide 47
Press TIA – Cat 6A Permanent Link near the bottom of the list, to select the TIA – Cat 6A Permanent Link test
Getting Started – Test Settings
Press

## Slide 48
TIA – Cat 6A Permanent Link
Getting Started – Test Settings
Press TIA – Cat 6A Permanent Link near the bottom of the list, to select the TIA – Cat 6A Permanent Link test

## Slide 49
Press Cable to select the type of cable to be tested
Getting Started – Test Settings
Press
Now we are going to select the cable manufacturer and cable type
The manufacturer, cable, and connector types are input for reporting purposes.

## Slide 50
Getting Started – Test Settings
Press
or Press
Cable
Scroll down and Press Berk-Tek

## Slide 51
Getting Started – Test Settings
Press
Cable
Scroll down and Press Berk-Tek

## Slide 52
Getting Started – Test Settings
Berk-Tek
Cable
Scroll down and Press Berk-Tek

## Slide 53
Getting Started – Test Settings
Press
Cable
Press LANmark-10G2

## Slide 54
LANmark-10G2
Getting Started – Test Settings
Cable
Let’s LANmark-10G2 for this test

## Slide 55
Getting Started – Test Settings
Press
Cable
Note that NVP is automatically set for this manufacturer for this cable
It can be edited, but let’s go with the pre-programmed setting
Now press Connector

## Slide 56
Cable
Scroll down and press Leviton
Getting Started – Test Settings
Press
or Press

## Slide 57
Cable
Scroll down and press Leviton
Getting Started – Test Settings
Press

## Slide 58
Cable
Scroll down and press Leviton
Getting Started – Test Settings
Press
Leviton

## Slide 59
Press eXtreme CAT 6A QuickPort
Getting Started – Test Settings
Press

## Slide 60
Press eXtreme CAT 6A QuickPort
Getting Started – Test Settings
Press
eXtreme CAT 6A QuickPort

## Slide 61
Getting Started – Test Settings

## Slide 62
Done with Test Settings
Let’s have a quick look at the Advanced Test Options
Press Test Options
(These are advanced features. We will show you these, but it’s ok to leave the defaults as is.)
Getting Started – Test Settings
Press

## Slide 63
Test Options. Here you can…
Locator
  Enable this to locate and display the worst NEXT and Return Loss (RL) fault in an Autotest result. Default is off.
Cable pairing type
  Select either T568-A or T568-B depending on cable pairing scheme
Autotest on connection
  If enabled (1), Autotest will automatically begin as soon as the connection between local and remote units is established. Default is off.
Direct Attach
  Enable this only if you are performing installation tests for Direct Attach Cables. Default is off
Enable AC Wiremap
  Enable AC Wire map to bypass testing for DC parameters. Default is off.
Press
Getting Started – Test Settings

## Slide 64
Test Options. Here you can…
Locator
  Enable this to locate and display the worst NEXT and Return Loss (RL) fault in an Autotest result. Default is off.
Cable pairing type
  Select either T568-A or T568-B depending on cable pairing scheme
Autotest on connection
  If enabled (1), Autotest will automatically begin as soon as the connection between local and remote units is established. Default is off.
Direct Attach
  Enable this only if you are performing installation tests for Direct Attach Cables. Default is off
Enable AC Wiremap
  Enable AC Wire map to bypass testing for DC parameters. Default is off.
Press
Getting Started – Test Settings

## Slide 65
Test Options. Here you can…
Locator
  Enable this to locate and display the worst NEXT and Return Loss (RL) fault in an Autotest result. Default is off.
Cable pairing type
  Select either T568-A or T568-B depending on cable pairing scheme
Autotest on connection
  If enabled (1), Autotest will automatically begin as soon as the connection between local and remote units is established. Default is off.
Direct Attach
  Enable this only if you are performing installation tests for Direct Attach Cables. Default is off
Enable AC Wiremap
  Enable AC Wire map to bypass testing for DC parameters. Default is off.
Press
Getting Started – Test Settings

## Slide 66
Test Options. Here you can…
Locator
  Enable this to locate and display the worst NEXT and Return Loss (RL) fault in an Autotest result. Default is off.
Cable pairing type
  Select either T568-A or T568-B depending on cable pairing scheme
Autotest on connection
  If enabled (1), Autotest will automatically begin as soon as the connection between local and remote units is established. Default is off.
Direct Attach
  Enable this only if you are performing installation tests for Direct Attach Cables. Default is off.
Enable AC Wiremap
  Enable AC Wire map to bypass testing for DC parameters. Default is off.
Press
Getting Started – Test Settings

## Slide 67
Test Options. Here you can…
Locator
  Enable this to locate and display the worst NEXT and Return Loss (RL) fault in an Autotest result. Default is off.
Cable pairing type
  Select either T568-A or T568-B depending on cable pairing scheme
Autotest on connection
  If enabled (1), Autotest will automatically begin as soon as the connection between local and remote units is established. Default is off.
Direct Attach
  Enable this only if you are performing installation tests for Direct Attach Cables. Default is off
Enable AC Wiremap
  Enable AC Wire map to bypass testing for DC parameters. Default is off.
Press
Getting Started – Test Settings
Press

## Slide 68
Press
Getting Started – Test Settings
Press

## Slide 69
Done with settings.
Now your ready to perform a Set Reference Test.
Getting Started – Test Settings

## Slide 70
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest
First, let’s look at the hardware
Getting Started – Perform set reference test

## Slide 71
What it is:
Perform a set reference test to eliminate losses in test cable from impacting test results
Why you need it:
When you test the installed cable at a facility or building, you don’t want to introduce additional errors when you bring your test cables and certifier. Therefore, before you begin testing the installed cable, first you want to eliminate losses from your certification tools. This is done through a procedure called “Set Reference Test”.
First, let’s look at the hardware
Getting Started – Perform set reference test
Eliminate losses due to these…

## Slide 72
Install Permanent Link Adapter in the Local Unit
Install Channel Adapter in the Remote Unit
Getting Started – Perform set reference test
Turn power on
Note: Permanent Link can also be installed in the remote unit, and Channel Adapter can be installed in the local unit.

## Slide 73
Plug in cable to Permanent Link Adapter in the Local Unit
Plug in cable to Channel Adapter in the Remote Unit
Check that units are connected
(red phone on each unit)
Cable included in kit
Getting Started – Perform set reference test

## Slide 74
Press the Tools One Touch button on the Local Unit
Local
Getting Started – Perform set reference test
Press

## Slide 75
Press Set Reference
Set reference must be performed if Local and Remote units are being paired for the first time
Perform set reference when starting a project
Getting Started – Perform set reference test
Press
Notes: Autotest will not run until a set reference is performed (or if there is a firmware mismatch)

## Slide 76
Note the message says to connect the local to the remote with a channel and permanent link probe
These are the probes, and this step has already been done a few slides ago when we verified the units were connected
Link probe
Channel probe
Press to start set reference
Getting Started – Perform set reference test
Press

## Slide 77
Test takes about 20 seconds to complete…
Getting Started – Perform set reference test

## Slide 78
Note that both the local and remote units provide test status
In this case, we have a message that says, “Set Reference Success”
Set reference is done
Now you are ready to perform Autotest – your first cable test!
Getting Started – Perform set reference test
See this video for more info on performing a Set Reference Test

## Slide 79
Getting started
Personalize your system settings
  Sound on/off, English/Metric, auto power off interval
Setup a project
  Site, user name, create cable labeling scheme or select your pre-configured List-based testing scheme create with the PC software
Setup your test to be performed
  TIA CAT 5e, CAT 6, CAT 6A…
Perform a set reference test to eliminate losses in test cable from impacting test results
Run your first test using Autotest
First, let’s look at the Adapter or Probes to be used
Getting Started – Run Autotest

## Slide 80
What it is:
With OneTouch Autotest, the cable test starts with all of your presets
With OneTouch Autotest, the cable label increments to the next cable, or with List-based testing, you can select the next cable from a list
When the test passes, it is automatically saved
Why you need it:
This is one of the most powerful tools of the WireXpert to help you get tests done fast
First, let’s look at the Adapter or Probes to be used
Getting Started – Run Autotest

## Slide 81
You will need to insert either…
…two Channel Adapters (channel probes)
Use these when performing channel tests with patch cords
…or two Permanent Link Adapters (Permanent Link Probes)
Use these when performing permanent link tests with reference cables
Hardware Required
Getting Started – Run Autotest

## Slide 82
Permanent Link Test
  Bring your test reference cables
  Perform set reference test
  Install Permanent Link Adapters and Test Reference Cords
  Plug your (male) reference cables into the customer’s (female) information outlets
Getting Started – Run Autotest

## Slide 83
Getting Started – Run Autotest
Channel Test
  Don’t need reference cables; use customer’s patch cord
  Perform set reference test
  Connect your WireXpert (female) Adapter to customer’s (male) patch cord
  Plug your (male) customer patch cable into the customer’s (female) information outlets

## Slide 84
Getting Started – Run Autotest
Press Autotest
Press

## Slide 85
Performing test…
Getting Started – Run Autotest

## Slide 86
You are standing at the remote
The “open” is 0.0 m from you; essentially, right in front of you at the connector
You are talking with the local side, and inform him you are going to fix the connector
The local guy says, “Great. I’ll wait. Just hit Autotest on your remote side when you are ready.”
Two displays, two processors, in both the local and the remote is called the Dual Control System™ which is powerful for getting testing done fast!
Click to see why
The remote unit displays a wire map
The remote unit says pair 78 is open at 0.0 m
In this case, the cable fails the test
5.0m
0.0m
0.0m
The wiremap shows conductor 8 is open

## Slide 87
Now, let’s say that you just fixed the connector, and it’s time to retest the repaired cable
Press Autotest on the remote side
5.0m
0.0m
Press

## Slide 88
Performing test…
Getting Started – Run Autotest

## Slide 89
Pass! Your first Autotest is complete
Congratulations! You just completed the getting started guide including…
Personalizing the settings on your certifier
Defining your project settings
Defining your test settings
Performing a Set Reference Test
And performing your first Autotest
Return to the Getting Started Guide
Return to the Main Menu

## Slide 90
Front Panel
Buttons and Connections
Resistive Touch Screen
Four one touch access buttons
For quick navigation
Automatic brightness
adjustments
Power Button
Navigation buttons for up-down, right-left navigation
Hanging cleats
Contact information for Softing
Connections
Select any of the functions for an overview

## Slide 91
Front Panel
Certifier Overview – Connections
Connections
Side of the Certifier
ActiveSync mini-USB port
- For use with eXport software to simulcast certifier screen on PC screen
3.5mm headphone jack
- So local operator can talk to remote operator
USB 2.0 for memory stick - upload reports
RJ-45 jack (future)
AC power charging port
Side of the Certifier

## Slide 92
Front Panel
Certifier Overview – Connections
Battery removal bottom of unit

## Slide 93
Front Panel
Certifier Overview – Brightness Adjustment
Automatic brightness
adjustments
Brightness adjustment has 4 pre-set brightness levels
Diming the screen can allow WireXpert to run longer on a charge

## Slide 94
Front Panel
Certifier Overview – Navigation Buttons
Navigation buttons for up-down, right-left navigation
Navigation buttons allow scrolling in long lists

## Slide 95
Front Panel
Certifier Overview – Hanging Cleats
Hanging cleats
Use hanging cleat or table top stand for conveniently positioning the WireXpert while working in the field so you can view the display while working with your hands
Table top stand

## Slide 96
Front Panel
Certifier Overview – Resistive Touch Screen
Resistive Touch Screen
Touch screen allows for easy navigation, even with gloves
Stylus provides an alternate method to navigate screens for your convenience

## Slide 97
Front Panel
Certifier Overview – One Touch Access Buttons
Autotest – One touch access to start next test
Setup – One touch access to settings
Data – One touch access to stored tests
Tools – One touch access to set reference, wire map
Four one touch access buttons
For quick navigation

## Slide 98
Why a OneTouch Data button?
This is OneTouch access to all of your stored cable certification test results
Find cables tested, both passed and failed
  Wiremap
  Length and Delay
  Resistance
  Insertion loss
  Return loss
  NEXT (near end crosstalk) value
  And more
Push the Data button
Press

## Slide 99
We are going to look back at the “LSU Tiger Stadium” test results
By pressing the current project called West Towne Mall, we can go back to the complete list of projects
Select West Towne Mall
Press

## Slide 100
By selecting “LSU Tiger Stadium” site, we will view test results stored for this project name
Select LSU Tiger Stadium
Press

## Slide 101
By selecting “LSU Tiger Stadium” site, we will view test results stored for this project name
Select LSU Tiger Stadium

## Slide 102
There are 19 tests that have passed, and 1 test that has failed in this project folder.
Press the magnifying glass to see these tests.
Press the magnifying glass
Press

## Slide 103
The local unit is showing a list of tests that have passed and failed.
Click any test to view results.
Press LSU Tiger StadiumLevel1TR1A12
Press

## Slide 104
Press the magnifying glass to see test results.
Press the magnifying glass
Press

## Slide 105
You can see…
The test passed
Wiremap
Key test parameters
You can also drill into these screens for more detail
Let’s click on “NEXT” to see near end crosstalk graph. Later we will come back to this screen and click the magnifying glass to view more test parameters.
Press NEXT test result
Press

## Slide 106
You can see…
The test passed
Wiremap
Key test parameters
You can also drill into these screens for more detail
Let’s click on “NEXT” to see near end crosstalk graph. Later we will come back to this screen and click the magnifying glass to view more test parameters.
Press NEXT test result

## Slide 107
This is the NEXT, or near end crosstalk, plot.
It shows a red test limit line…

## Slide 108
This is the NEXT, or near end crosstalk, plot.
It shows a red test limit line…
…and this test passes because the amplitude results from the cable test are below the red test limit line.

## Slide 109
This is the NEXT, or near end crosstalk, plot.
It shows a red test limit line…
…and this test passes because the amplitude results from the cable test are below the red test limit line.
The “margin” is the area under the curve between the red test limit line and the test results.

## Slide 110
This is the NEXT, or near end crosstalk, plot.
It shows a red test limit line…
…and this test passes because the amplitude results from the cable test are below the red test limit line.
The “margin” is the area under the curve between the red test limit line and the test results.
If the data from the test exceeds the test limit line, a test will fail. Higher margin means that the cable passes with more head room.
Let’s go back one screen and view more test parameters.
Press
Press

## Slide 111
Press the magnifying glass to see more test parameters
Press
Press

## Slide 112
This is a complete list of test parameters and values.
List includes…
Many more test results are available, which can be useful when you are in the field and troubleshooting a failed test.
Return to overview
Return to menu
Notes: RL – return loss locator…
If you have a bad end, you look at the RL data screen, where ever you have the bigger spike, that’s the end you want to start troubleshooting. If you have spikes in the middle, you have a problem in the middle of the cable. You can turn pairs on and off, so you can isolate pairs, then you can find the pair that causes the problem, and then terminate the one problem pair/connection.

## Slide 113
Front Panel
Certifier Overview – Contact Information
Contact information for Softing
Press the for list of Softing contacts
Press tools, and about us, for list of Softing contacts
Sales: +1.865.251.5252
sales@softing.us
Support +1.865.251.5244
support@softing.us

## Slide 114
Press here to turn on power

## Slide 115
Booting…

## Slide 116
Booting…

## Slide 117
Boot Ver 2.3
Booting…

## Slide 118
OS Ver 7.2
Booting…

## Slide 119
OS Ver 7.2
Booting…

## Slide 120
4500
Booting…

## Slide 121
4500
Booting…

## Slide 122
4500
Booting…

## Slide 123
4500
Booting…

## Slide 124
Ready
Boot complete!
Click here to see Power LED color and charging overview
Click here to return to WireXpert Overview

## Slide 125
Plug-in unit to charge
Green means not plugged-in, not charging, and (just) on

## Slide 126
Charging
Yellow/red means plugged-in, charging, and on

## Slide 127
Charging
Red means plugged-in charging, and off
A fully charged battery can operate WireXpert for 5 to 8 hours

## Slide 128
Charged
LED off means plugged-in completely charged, and power off
A fully charged battery can operate WireXpert for 5 to 8 hours

## Slide 129
Charged
Green LED, plugged-in means completely charged, and on
A fully charged battery can operate WireXpert for 5 to 8 hours

## Slide 130
What if you want to know the remaining hours you have left on the battery in the local or the remote units?
Click here for a secret menu item…
Charging
Charging icon or percent charge

## Slide 131
Charging
Click the space between the phone icon and the charge icon

## Slide 132
Charging
Both local and remote units are charging

## Slide 133
Charging
Local unit has 47 minutes remaining
Remote unit is charging

## Slide 134
Charging
Local unit has 47 minutes remaining
Remote unit has 7 hours and 15 minutes remaining
Hours remaining can be very helpful when trying to figure out if you have enough battery life to finish the job
Return to Overview

## Slide 135
Why an Autotest button? When you are ready to start certifying cables, press the Autotext button…
  The test will start
  The test will use your pre-sets
This is one of the most powerful tools of the WireXpert to help you get tests done fast!
Autotest
Autotest

## Slide 136
Now, you’ve fixed the connector, and it’s time to retest the repaired cable
Press Autotest on the remote side
Autotest
Press

## Slide 137
Performing test…
Autotest

## Slide 138
Pass! Your first Autotest is complete
Return to OneTouch Access
Return to Menu or Home
Autotest

## Slide 139
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
System Settings
  Sound on/off, English/Metric, auto power off interval, audio volume
  Click Next to go to Quick Setup
4
3
2
1

## Slide 140
Press the Quick Setup Menu item on the Certifier
Press
Setup
Quick Setup
  Select your test and go!
Why Quick Setup?
If you have already setup your System Settings
  Feet/meters
  Plot axis preferences
  Audio settings
And, you have already defined your Project Settings
  Site
  User
  Cable, connector, shield
Then just go to Quick Setup, select the test standard (here), and hit Autotest
This is a super fast way to get certification testing started fast

## Slide 141
Select CAT 6A standard
Quick Setup
Press
Setup
Quick Setup
  Select your test and go!

## Slide 142
Quick Setup
That’s it!
Setup is that fast, and now you’re ready to press Autotest
Press Autotest or click next to move on to Test Settings
Press
Setup
Quick Setup
  Select your test and go!

## Slide 143
Performing test…
Autotest

## Slide 144
0.00m
Pass! Autotest is complete
Autotest
6A Permanent
6A Permanent

## Slide 145
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Press
Why Test Settings?
When you certify cables, you will be asked to certify the cable to a specific standard, or you are performing warranty work
Standards such as a TIA CAT 6 limit standard or an ISO limit standard can be selected here
It is likely you will access Test Settings before starting a job
You will access Test Settings before performing a “Set Reference” or “Autotest”
Set Reference and Autotest are covered in the getting started guide from the main menu
Press Test Settings

## Slide 146
From this screen, you can select the…
Test standard
Cable being tested
Cable manufacturer
Cable connector
Note: this is powerful because the NVP value is automatically loaded for you, and you will have a record of the cable manufacturer, cable type, and connector for future reference
Getting Started – Setup a test
Press
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Press Test Limit

## Slide 147
Getting Started – Setup a test
Press
The local unit is now showing a list of limit standard categories
Press TIA Standard
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…

## Slide 148
Getting Started – Setup a test
The local unit is now showing a list of TIA standards
Click next to move on to Project Settings
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…

## Slide 149
Press
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
Why Project Settings?
This is like creating a folder for your Autotest results
When a user presses the OneTouch Autotest button (Autotest covered later) and runs a cable test, the results of the cable test will be placed in this user defined folder
Creating a user name, and a site folder is a great way to organize data for a given job site
Press the Project Settings Menu item on the Certifier

## Slide 150
Getting Started – Setup a project
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
Project settings include creating a
Site name
User name (operator)
Selecting a label source
Designating an autosave preference
Click Next to move on to System Settings

## Slide 151
Press the System Settings Menu item on the Certifier
Press
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
System Settings
  Sound on/off, English/Metric, auto power off interval, audio volume
Why system settings? These are your preferences like engineering units in feet or meters, for example. Once set, you will not need to access these very often.

## Slide 152
Getting Started – Personalize your system settings
Return to OneTouch Access
Return to Menu or Home
Note some of the settings in the menu on the Local Unit
Setting cable length in units of feet or meters
Plot y-axis preferences
Device sleep interval
Auto power off interval
Audio volume preferences
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
System Settings
  Sound on/off, English/Metric, auto power off interval, audio volume
Done with Setup!

## Slide 153
Getting Started – Personalize your system settings
Return to OneTouch Access
Return to Menu or Home
Setup
Quick Setup
  Select your test and go!
Test Settings
  TIA CAT 5e, CAT 6, CAT 6A…
Project Settings
  Site, user name, create cable labeling scheme
System Settings
  Sound on/off, English/Metric, auto power off interval, audio volume

## Slide 154
Press the Tools
Press

## Slide 155
Why tools? Here you will find useful pre-test tools, and troubleshooting tools. For example…
Set Reference
  Used to subtract losses due to the test reference cables
  Use this tool anytime you start a project for the first time
Length & Delay
  Used to determine cable length
Wiremap
  Used to verify that the wire is mapped from one end of the cable to the other end of the cable correctly
  Wiremap will detect opens and will tell you if wires are terminated to the wrong pin
  Wiremap can be used to test, but not certify, cables longer than test standard limits will allow
Learn NVP
  Most of the time, you have the NVP from the manufacturer, but in some cases, you may not have this value, so the WireXpert can determine this for you
Reset Probe
  The probe in this case is the adapter or the small, blue removable attachment on the back of each WireXpert
About us
  Provides you with contact information back to Softing support
Factory Calibration
  This button is for Softing to use during calibration
Set Reference Test

## Slide 156
Install Permanent Link Adapter in the Local Unit
Install Channel Adapter in the Remote Unit
Set Reference Test – Adapters Required

## Slide 157
Plug in cable to Permanent Link Adapter in the Local Unit
Plug in cable to Channel Adapter in the Remote Unit
Check that units are connected
Cable included in kit
Local
Set Reference Test – Cables Required
Video: how to connect cables

## Slide 158
Press the Tools One Touch button on the Local Unit
Local
Set Reference Test – Cables Required
Press

## Slide 159
Set Reference Test – Press Set Reference
Press Set Reference
Press

## Slide 160
Note the message says to connect the local to the remote with a channel and permanent link probe
These are the probes, and this step has already been done.
Link probe
Channel probe
Press to start set reference
Set Reference Test – Start Test
Press

## Slide 161
Test takes about 20 seconds to complete…
Set Reference Test – Test Started

## Slide 162
Set reference is done! Success!
Note that both the local and remote units provide test status
Set Reference Test – Success!
See this video for more info on performing a Set Reference Test
Return to OneTouch buttons overview
Return to Menu or Home

## Slide 163
Support, XpertCare
Key benefits of XpertCare
Covers accidental damage of WireXpert local or remote Dual Control System™
Shipping costs are covered for WireXpert repair
Loaner units provided next business day for WireXpert repair so your job can keep moving
Annual calibration provided
  Loaner units provided when scheduled at least 2 weeks in advance for WireXpert calibration
Permanent link test reference cables covered for manufacturing defects, and once per year for accidental damage
Adapters covered for manufacturing defects
Batteries covered for manufacturing defects

## Slide 164
Support, XpertCare
Why XpertCare?
Protect your investment
Keep jobs going when accidents happen
Description | Cost
Unit Repair | $895
Loaner unit | $350
Annual calibration | $480
Accessory replacement | $500
Total Separate Services | $2,225
XpertCare 1-yr package | $1,070
Savings: 52% | $1,155

## Slide 165
eXport Software
This is a list of tests for a named project

## Slide 166
eXport Software
Note that if you right click a cable label, a menu of convenient options appears
Since the WireXpert stores raw test data, you can even retest, or re-certify, a cable to a different standard just by selecting a different standard and selecting your stored data results

## Slide 167
eXport Software
Let’s pick a test and drill-in to see results
Select Test 6
Select

## Slide 169
Select any tab
Note: the red x on the tab quickly points you to the failed test parameter

## Slide 170
Select any tab

## Slide 171
Select any tab

## Slide 172
Select any tab
Test limit exceeded

## Slide 173
Select any tab

## Slide 174
Select any tab

## Slide 175
Select any tab

## Slide 176
Select any tab

## Slide 177
Select any tab

## Slide 178
Select any tab

## Slide 179
Select any tab

## Slide 180
Select any tab